## Slide 1: Structural and chemical characterization of epitaxial layers
P. Vennéguès
Centre de Recherche sur l’Hétéroépitaxie et ses Applications
UMR7073
Sophia Antipolis

## Slide 2: Structural and chemical characterization of epitaxial layers
P. Vennéguès
Centre de Recherche sur l’Hétéroépitaxie et ses Applications
UMR7073
Sophia Antipolis

## Slide 3
Introduction
CRHEA:
      8 molecular beam epitaxy reactors
      6 Metalorganic vapor phase epitaxy reactors
      Studied materials:
        III-N (GaN, AlN , InN, ….)
        ZnO-based
        SiC
        2D: graphene, TMDs…
        halogenated perovskites (Cs(Pb/Sn)Br type)
        II-N
CAT team (Advanced characterization and Technology)

## Slide 4
Introduction
Ex-situ characterization of heteroepitaxial layers by « routine » technics accessible in a material lab
2 levels: « basic » characterizations ; « advanced » characterizations
Examples of characterizations mainly based on materials elaborated in CRHEA

## Slide 5
Plan
Atomic force microscopy (AFM)
Scanning electron microscopy (SEM)
X-Ray diffraction (XRD)
Transmission electron microscopy (TEM)

## Slide 6
Introduction
Ix1y1()
x
y
x1
y1
sample
image
Ix1y1()
y1
x1
Scanning probe microscopies:
AFM
SEM
STEM
sample
Fixed probe
image
spectrum
Fixed probe:
XRD
TEM

## Slide 7
Atomic force microscopy
https://culturesciences.chimie.ens.fr/thematiques/chimie-analytique/la-microscopie-a-force-atomique-pour-l-observation-de-molecules-avec
https://www.afmworkshop.com/
https://chem.libretexts.org/Bookshelves/Analytical_Chemistry/Supplemental_Modules_(Analytical_Chemistry)/Microscopy/Scanning_Probe_Microscopy/03_Basic_Theory/02_Atomic_Force_Microscopy_(AFM)
Si or SiN tip
10-15µm
Apex qq nm
A. Courville: aimeric.courville@crhea.cnrs.fr
F. Bartoli: florian.bartoli@crhea.cnrs.fr

## Slide 8
Atomic force microscopy
Contact mode:
Uses repulsive forces between tip and sample
Piezoelectric element maintains constant tip-to-sample distance
Sensitive to surface contamination: loss of resolution
Possible damage to tip

## Slide 9
Atomic force microscopy
Tapping mode (tapping):
Uses repulsive forces between tip and sample
Periodic vibration of tip-sample distance
Variation of oscillation amplitude: image
High resolution
Preserved tip

## Slide 10
Atomic force microscopy
Piezoelectric module:
X-Y displacement (scanning)
Z displacement

## Slide 11
Atomic force microscopy
Calibration Si grid:
(X,Y) = 10000 mm2; Z = 0,25 mm
GaN quantum dots
(X,Y) = 0,25 mm2; Z = 5 nm
X = 100 mm
Y = 100 mm
Z = 250 nm
X = 0,5 mm
Y = 0,5 mm
Z = 5 nm
Field of view: 100µmx100µm to 0,5µmx0,5µm
Résolution: lateral, a few nm; vertical, a few 10pm
a few 10pm

## Slide 12
Atomic force microscopy
Possible artefacts for 3D objects

## Slide 13
Atomic force microscopy
Rms: 2nm
Rms: 1,2nm
Face to face annealing of a AlN/BN/sapphire
A. Zaiter et al, Materials 2022, 15, 8602.
“basic” use of AFM to optimize the elaboration of heteroepitaxial layers

## Slide 14
Atomic force microscopy
Determination of dislocation density in III-N
Example of GaN MOVPE on sapphire
dislocation
line tension
surface tension
F.C. Frank, Acta Crystallogr. 4, 497 (1951)
r
r = (b2)/(82s)
b = Burger vector
= shear modulus
s = surface energy
2 main dislocation types in III-N:
a: line ≈ vertical; edge type; ba= 0.3189nm;
a+c: line ≈ vertical; mixed type; ba+c=0.5185nm

## Slide 15
Atomic force microscopy
a+c-type dislocations; screw component=>2 atomic planes terminations
a-type dislocations
?: dislocation with b>ba+c; 4 atomic planes terminations; b= 2c or b=a+2c
Determination of dislocation density in III-N
Example of GaN MOVPE on sapphire

## Slide 16
Atomic force microscopy
2µm
1’ annealing at 1060°C under SiH4 et NH3
R.A. Oliver et al. 2006 J. Cryst. Growth. 289, 506
M. Khoury et al., Semicond. Sci. Technol. 28 (2013) 035006
Determination of dislocation density in III-N
Example of GaN MOVPE on sapphire

## Slide 17
Atomic force microscopy
Quantum dots GaN/AlN/ sapphire
Height: 12nm, width: 32nm

## Slide 18
Atomic force microscopy
Different stages of the growth of Graphene on SiC
thèse Chiara Mastropasqua U. Côte d’azur
C. Mastropasqua et al., npj 2D Materials and Applications | (2025) 9:32

## Slide 19
Atomic force microscopy
+
Easy to use: Relatively short training
Air observation
Fast
High resolution, especially in Z (qq 10 pm)
-
Possible artifacts: limit to very smooth samples
Information limited to surface topography

## Slide 20
Atomic force microscopy
GaN NH3-MBE
As-grown
15’ annealing at 1000°C under NH3

## Slide 21
Atomic force microscopy
Force Modulation (FM) Microscopy
Lateral Force Microscopy (LFM)
Force-distance (F vs d) Measurements
Conductive Atomic Force Microscopy (C-AFM)
Piezoresponse Force Microscopy (PFM)
Electrostatic Force Microscopy (EFM)
Scanning Capacitance Microscopy (SCM)
Scanning Spreading Resistance Microscopy (SSRM)
Magnetic Force Microscopy (MFM)
Kelvin Probe Force Microscopy (KPFM)
….
Measuring the elastic properties of surfaces
Required a conductive tip: metal-coated tip => loss of resolution

## Slide 22
Atomic force microscopy
Force Modulation (FM) Microscopy
Lateral Force Microscopy (LFM)
Force-distance (F vs d) Measurements
Conductive Atomic Force Microscopy (C-AFM)
Piezoresponse Force Microscopy (PFM)
Electrostatic Force Microscopy (EFM)
Scanning Capacitance Microscopy (SCM)
Scanning Spreading Resistance Microscopy (SSRM)
Magnetic Force Microscopy (MFM)
Kelvin Probe Force Microscopy (KPFM)
….
https://www.bruker.com/en/products-and-solutions/microscopes/materials-afm/resource-library/e-book-the-definitive-afm-modes-handbook.html
Required a conductive tip: metal-coated tip => loss of resolution
Measuring the elastic properties of surfaces

## Slide 23
Atomic force microscopy
C-AFM (conductive AFM)
AFM in contact mode
Simultaneous acquisition of a topographic image and an electrical conduction image
With a fixed probe, I(V) measurements can be taken by varying the applied voltage

## Slide 24
Atomic force microscopy
C-AFM (conductive AFM)
https://www.parksystems.com/en/learning-center/lc-detail.learning875

## Slide 25
Atomic force microscopy
PFM (Piezoresponse force microscopy)
Contact mode
Mapping of polarity domains
Polarity can be reversed by applying a sufficiently large field (ferroelectric materials)

## Slide 26
Atomic force microscopy
PFM (Piezoresponse force microscopy)
-6V
+6V
-6V
Ferroelectric effect in HfO2
Initial step

## Slide 27
Scanning Electron Microscopy
https://www.globalsino.com/EM/; https://myscope.training/SEM_SEM_Basics
x
y
x1
y1
sample
image
Ix1y1()
Probe: focus electron beam=> HT (500meV to 30kV) and I can be tuned
SE: topography contrast
BSE: chemical contrast
X-rays: composition maps
CL: optical properties map
V. Brandli: virginie.brandli@crhea.cnrs.fr

## Slide 28
Emitter stage
Emission and acceleration of e-
Condenser stage
Shaping the beam
Objective stage
Focusing the beam
Scanning the beam
Sample stage
Sample and detectors
Scanning Electron Microscopy
Under vacuum

## Slide 29
Scanning Electron Microscopy
Non destructive technique
Possibility to analyze large and thick samples (full wafer)
large range of magnification: x10 to x2M
Tilt and rotation of the samples
No sample preparation required except:
  For insulating samples: coating with Au or C
  For cross section observations (layers thickness measurement)
Under vacuum

## Slide 30
Emitter
Scanning Electron Microscopy
Thermionic emitters: W or LaB6 filaments
Field emission gun emitters:
  Cold FEG: high brightness, small beam
  Schottky FEG: high current, more versatile
Table top SEM

## Slide 31
Scanning Electron Microscopy
Interaction volume
Depends on HT and material
Elastic interactions:
BSE: chemical contrast; resolution defined by probe size and interaction volume: 10-100nm
Inelastic interactions:
SE
X-rays
Optical emission (CL)
Si
AlN
AlGaN1
AlGaN2
AlGaN3

## Slide 32
Scanning Electron Microscopy
Secondary electrons
<50V
Comes from the top surface of the specimen (a few nm)
Contrast related to the surface topography
Resolution defined by probe size and interaction volume:
    HT=> probe size: improvement of resolution
    HT => interaction volume: deterioration of resolution
    HT => more BSE detected in the SE detector: use of in-lens detector
Resolution: trade off between different parameters and is sample dependent; down to <1nm

## Slide 33
Scanning Electron Microscopy
Low WD (a few mm):
highly convergent beam: high resolution
Limited depth of field
Problem with 3D samples; tilting ….
Less space for detectors (EDS, CL…)
Large WD (<10mm):
More parallel beam: limited resolution
High depth of field
Possibility to observe 3D samples and to tilt the samples
Introduction of different kinds of detectors

## Slide 34
Scanning Electron Microscopy
small WD
GaN QD/AlN

## Slide 35
Scanning Electron Microscopy
Large WD
SEM very well adapted to the study of 3D µ- and nano-objects

## Slide 36
Scanning Electron Microscopy
AlGaN µ-object on sapphire
Charging problems:
  Coating with conductive layer (C, Au…)
  Low HT
  Low vacuum (small gas flow on the sample surface)

## Slide 37
Scanning Electron Microscopy
X-Rays: when an incident e- ionized an atom, it returns to its ground state by emitting characteristic X Rays
Wavelength-dispersive spectrometer (WDS): highly sensitive but requires a dedicated instrument
Energy-sensitive spectrometer (EDS or EDX): solid state detector capable of analyzing X-ray energies
AlGaScN

## Slide 38
Scanning Electron Microscopy
Simultaneous e- image and EDS maps of a semiconductor device in cross section
EDS
Bruker

## Slide 39
Scanning Electron Microscopy
EDS quantitative analysis: use of standards (sample with same elements of known concentration)
Semiquantitative analysis: comparison with standards from a library
Quantitative analysis: standards acquired with the same microscope and the same acquisition parameters; corrections to be applied:
  Z: corrections relying the mean Z number of the sample
  A: absorption
  F: fluorescence
Quantitative composition determination by EDS is not straightforward
a few % precision
Not usable for light elements (Z<11)

## Slide 40
Scanning Electron Microscopy
Quantitative EDX for adjusting growth parameters of alloys (AlGaN)
substrate
layer
Decrease HT
substrate
layer
Increase layer thickness
Casino sotfware

## Slide 41
Scanning Electron Microscopy
BSE may be diffracted by the crystalline lattice of the sample
e- beam
Diffraction cones
Kikuchi bands
Kikuchi pattern
Analysis of Kikuchi patterns and comparison to calculated patterns allow obtaining:
  Phase map
  Orientation map
Popular for polycrystalline materials (metals) but thanks to improvement of resolution (spatial ≈ nm and angle <0.1°) EBSD may be useful for heteroepitaxial layers
Electron backscattered diffraction
EBSD
https://mea.edu.umontpellier.fr/ebsd/ebsd-principe/

## Slide 42
Scanning Electron Microscopy
F.A. Ponce Group III Nitride Semiconductor Compounds, Physics and Applications (Clarendon Press, Oxford 1998).
EBSD
III-N heteroepitaxial microstructures
Dislocations in the domains boundaries to compensate the misorientations
R. Dagher et al, J. Cryst. Growth, 526, 2019, 525235
K. Baril et al., J. Appl. Phys., 133, 245702, (2023)
SiO2
GaN

## Slide 43
Scanning Electron Microscopy
Reduction of misorientations between adjacent pillars forming different groups
EBSD
K. Baril, PHD U. Côte d’Azur
M. Wehbe et al., J. Appl. Phys. 136, 224302 (2024)

## Slide 44
Scanning Electron Microscopy
sonde
Electron channeling contrast imaging ECCI
Revealing crystal imperfection thanks to local variations of diffraction conditions (see TEM chapter)
G. Naresh-Kumar et al., Materials Science in Semiconductor Processing, 47(2016)44–50

## Slide 45
Scanning Electron Microscopy
ECCI
PHD Audrey Gilbert, U. Montpellier
Gradual relaxation of AlSb on GaSb with the introduction of misfit dislocations (MBE)

## Slide 46
Scanning Electron Microscopy
Cathodoluminescence CL
Advantages as compared to PL:
Study of wide bandgap materials (Al-rich AlGaN alloys; AlN bandgap≈ 6.2eV )
Spatial resolution
CL modes:
  panchromatic: using all the emitted light to create images
  Hyper spectral analysis

## Slide 47
Scanning Electron Microscopy
CL
Panchromatic image
MOVPE GaN
Dislocation are non-radiative carriers recombination centers => they appear as dark spots on a emitting (bright) matrix
Useful to determine dislocations densities in high crystalline quality layers
M. Khoury et al. Semicond. Sci. Technol. 28 (2013) 035006
K. Baril, PHD U. Côte d’Azur

## Slide 48
Scanning Electron Microscopy
CL
Hyperspectral analysis
P.M. Coulon et al., J. Appl. Phys. 115, 153504 (2014)
MOVPE GaN µwire on sapphire
a: N
b
A: Ga
B
Ga-polarity
N-polarity
TEM image

## Slide 49
X-Ray diffraction
Maud Nemoz: maud.Nemoz@crhea.cnrs.fr
image
spectrum
sample
Fixed probe
No lens for Xray: no image
Typical beam size with a laboratory tool: 1cmx50µm
Typical penetration length: a few 10µm
average information on a large volume of sample
Non destructive technique
Maud Nemoz

## Slide 50
X-Ray diffraction
Xray diffraction: study of the reciprocal space
GaN [11-20]
d(0002)
1/d(0002)
(0002)
(000-2)
d(1-100)
1/d(1-100)
(1-100)
(-1100)
(0002)
(0004)
(0000)
(1-100)
(1-101)
(1-102)
(1-103)
(1-104)
Direct space
Reciprocal space

## Slide 51
X-Ray diffraction
Xray diffraction: study of the reciprocal space based on the Bragg law
2d(hkl) sin  = n with d(hkl) interreticular distance of hkl plane,  diffraction angle and  Xray wavelength
Q = kf – ki Q reciprocal space vector, ki incident wave vector, kf final wave vector
|Q | = 1/d, |k | = 1/ 
sample
Fixed probe
sample
Fixed probe
Study of the 3D reciprocal space by varying the sample/Xray beam orientation

## Slide 52
X-Ray diffraction
Xray source:
  Xray tube (e-beam hit a target which emit characteristic Xray)
  Rotating anode
Characteristic emission
Xray source:
  Synchrotron radiation
Anode | Kα1(Å)
Cu | 1.54060
Cr | 2.28970
Fe | 1.93604
Co | 1.78897
Mo | 0.70930

## Slide 53
X-Ray diffraction
diffracted beam capture
-slits
-analyzer
-detector
Sample stage
Xray source
Slit
monochromator
mirror
Incident beam condition

## Slide 54
X-Ray diffraction
To obtain a small divergence into the diffraction’s plane
Only the Kα1 can go out the monochromator
Xray source
(0D, 1D, 2D)

## Slide 55
X-Ray diffraction
4 axes goniometer: allow to access any position in the 3D reciprocal space
φ

## Slide 56
X-Ray diffraction
K. Inaba et al., Advances in Materials Physics and Chemistry, 2013, 3, 72-89
Diffraction geometries

## Slide 57
Qz
Qy
2
22
detector
1
21
Incident Xrays
kf
ki
X-Ray diffraction
/2 symmetric geometry: scan of the reciprocal direction perpendicular to the sample surface: study of the planes parallel to the surface

## Slide 58
X-Ray diffraction
Qz
Qy
Detector(fixed)
21
Incident Xrays
kd
ki

 scan symmetric geometry: fixed detector (2); determination of the diffraction peak width in the Qy direction (in plane)

## Slide 59
X-Ray diffraction
Detector
Qz
Qy

2
 = + 
Asymmetric geometry: study of inclined planes
  /2: scan of a direction inclined by  to the normal of the film
   scan
  Combining /2 and/or  scans: 2 D reciprocal space mapping (rsm)

## Slide 60
X-Ray diffraction
GaN (1-210)
symmetric
asymmetric
/2

rsm

## Slide 61
X-Ray diffraction
Study of epitaxial relationships: MBE-grown Mg3N2/MgO
(222)
Growth direction
(440)
Mg3N2 (1-10)
φ scans (440) planes

## Slide 62
X-Ray diffraction
Study of epitaxial relationships: MBE-grown Mg3N2/MgO
Presence of 12 (400) peaks instead of 3 expected
Presence of rotational domains
[1-10] Mg3N2//[001]MgO
[11-2] Mg3N2 //[01-1]MgO
Plus 90°, 180° and 270° in plane rotations
P. John et al., Phys. Rev. Mat. 04, 054601 (2020)
P. John et al. J. Appl. Phys. 129, 095303 (2021)

## Slide 63
X-Ray diffraction
Lattice parameters determination
ac
as
ac > as
dilatation
contraction
b
Coherent growth
Elastic relaxation
Incoherent growth
Plastic relaxation with introduction of misfit dislocations above a critical thickness

## Slide 64
X-Ray diffraction
Lattice parameters determination
GaN (1-210)
symmetric
 /2 : determination of out-of-plane lattice parameter
2d(hkl) sin  = n
AlN/GaN/SiC
SiC
GaN
AlN
The position of characteristic peaks in  /2 scan allow determining the strain state

## Slide 65
X-Ray diffraction
Lattice parameters determination
GaN (1-210)
Position of symmetric peaks ( /2) : function of out-of-plane lattice parameter
Position of asymmetric peaks ( /2): function of out-of-plane and in plane lattice parameters
Determination of in and out-of-plane latices parameters
Qz
Qy

## Slide 66
X-Ray diffraction
Lattice parameters determination
GaN (1-210)
rsm
Out-of-plane
In-plane

## Slide 67
X-Ray diffraction
Lattice parameters determination
GaN pillar
Al0.15Ga0.85N
Al0.83Ga0.17N
20µm
2µm
In-plane
Out-of-plane
Al0.83Ga0.17N and Al0.15Ga0.85N have same in-plane parameters; Al0.83Ga0.17N fully relaxed
Compressive strain of Al0.15Ga0.85N on Al0.83Ga0.17N: out-of-plane dilatation

## Slide 68
X-Ray diffraction
Crystalline quality
substrate
lateral
coherence
Length L
vertical
coherence
Length L
tilt angle
twist angle
Qz
Qy
f(1/L)
f(1/ L, tilt)
f(1/ L)
f(tilt)
Mosaic model for heteroepitaxial films microstructure
Qx
Qz = f(1/L)
Qy = f(1/ L, tilt) for symmetric peaks
Qx = f(1/ L, twist) for asymmetric peaks

## Slide 69
X-Ray diffraction
Crystalline quality
Qz
Qy
f(1/L)
f(1/ L, tilt)
f(1/ L)
f(tilt)
Qz = f(1/ L, tilt) for symmetric peaks
Qx = f(1/ L, twist) for asymmetric peaks
Tilt (twist) contribution to the peak width increases with Qz whereas 1/Lcontribution is constant => Williamson-Hall plot
fwhm(sin)/
(sin)/
Y = (tilt  x) + 0.9/2 L
(0002)
(0004)
(0006)
Typical Williamson-Hall plot
for a III-N layer (symmetric peaks)
Qx

## Slide 70
X-Ray diffraction
Crystalline quality: MBE-grown Mg3N2/MgO
Williamson-Hall « like » plot, asymmetric peaks:
Tilt: 1,3°
Twist: 2,5°
P. John et al. J. Appl. Phys. 129, 095303 (2021)
R. Lee et al., Appl. Phys. Lett. 86, 241904 (2005)

## Slide 71
X-Ray diffraction
Thickness measurement: symmetric geometry
Scherrer equation: t= λ/(Δ2θcos θ)
GaNlayer
1 µm
200 nm
50 nm
Main peak
Thickness fringes
GaN (0002)
Δ2θ

## Slide 72
X-Ray diffraction
Thickness measurement: reflectivity (MBE-grown Mg3N2/MgO)
From best fit:
Mg3N2= 98nm
MgO= 42nm
Interface roughness= 1nm
Mg3N2 density= 2.56 g cm−3

## Slide 73: Transmission Electron Microscopy
I. Florea
Transmission electron microscopy:
    Requires a 0 to a few hundred nm-thick sample => sample preparation
Cross section (CS)
Plan view (PV)
- “classical” preparation: cutting, gluing, polishing and ion thinning; both CS et PV; time consuming, manual skills, non localized, high quality very thin samples
- Focus ion beam

## Slide 74: Transmission Electron Microscopy
I. Florea
Transmission electron microscopy:
    Requires a 0 to a few hundred nm-thick sample
Cross section (CS)
Plan view (PV)
- “classical” preparation: cutting, gluing, polishing and ion thinning; both CS et PV; time consuming, manual skills, non localized, high quality very thin samples
- Focus ion beam: CS; quick (1/2 day/sample), partially automatic, localized; amorphous surface layer, expensive
Destructive technic

## Slide 75: Transmission Electron Microscopy
Transmission electron microscopy
Scanning transmission electron microscopy
sample
Fixed probe
image
“large” parallel beam on the sample
“small” convergent beam on the sample
x
y
x1
y1
sample
image
Ix1y1()

## Slide 76: Transmission Electron Microscopy
Transmission electron microscopy
Objective lens focus plane: diffraction
image

## Slide 77: Transmission Electron Microscopy
Electron diffraction
(0002)
(10-10)
GaN [11-20] zone axis
(0002)
(11-20)
GaN [10-10] zone axis
Check of the mono-crystallinity of the layer

## Slide 78: Transmission Electron Microscopy
Electron diffraction
Determination of epitaxial relationship
ZnO on R-plane: sapphire in black, ZnO in gray
[1−100]ZnO// [11−20]sap et [0001]ZnO// [−1101]sap
P. Vennéguès et al., J. Appl. Phys. , 103, 083525 (2008)
Polarity determination by convergent beam electron diffraction
83nm
40nm
Ga
N
c

## Slide 79: Transmission Electron Microscopy
TEM imaging
Diffraction contrast
0
g
-g
-g
-2g
0
Incident beam
sample
Objective lens
aperture
Bright field
Dark field

## Slide 80: Transmission Electron Microscopy
TEM imaging
Diffraction contrast
Diffraction contrast: defects characterization
No strain of horizontal planes: no contrast on DF using the corresponding diffraction spot g1
Strain on vertical planes: contrast on DF using the corresponding diffraction spot g2
Extinction rules for a defect with a displacement field R:
g.R= n integer => no contrast
g.R ≠ n => contrast

## Slide 81: Transmission Electron Microscopy
TEM imaging
Diffraction contrast: defects characterization
Extinction rules for a defect with a displacement field R
g.R= n integer => no contrast
g.R ≠ n => contrast
a
a+c
c
a = 1/3<11-20>
c = <0001>
a+c = 1/3<11-23>
Dislocations in III-N
g/b | a= 1/3 <11-20> | c= <0001> | a+c = 1/3 <11-23>
{11-20} | contrast | Out-of-contrast | contrast
{0002} | Out-of-contrast | contrast | contrast

## Slide 82: Transmission Electron Microscopy
TEM imaging
g/b | a= 1/3 <11-20> | c= <0001> | a+c = 1/3 <11-23>
{11-20} | contrast | Out-of-contrast | contrast
{0002} | Out-of-contrast | contrast | contrast
Close to <1-100> zone axis
MOVPE GaN on sapphire:
Most of dislocations are a+c type
2 a-type
(0002)
(11-20)
diffraction contrast
a
b
300nm
(0002)
(11-20)
a-type
a+c-type
c-type

## Slide 83: Transmission Electron Microscopy
High resolution-TEM imaging
0
g
-g
HR-TEM imaging: image simulation required
Thickness (nm)
GaN (11-20)

## Slide 84: Transmission Electron Microscopy
Scanning TEM

## Slide 85: Transmission Electron Microscopy
Scanning TEM
Annular detector 1
Annular detector 2
Adjustment of the detection angle playing with the diffraction magnification (camera length)
ADF: diffraction contrast
HAADF: Z contrast
ADF+HAADF
HAADF
High convergence angle: high resolution
I≈ Z2
HR-HAADF images directly interpretable

## Slide 86: Transmission Electron Microscopy
(Scanning) transmission electron microscope:
HT: 30 to 300kV; 200-300kV for material science; low HT for fragile materials (2D)
Spatial resolution:
  non corrected: depends on the objective lens: 2-2.5Å
  Aberration corrected (probe or objective): below 1Å
EDX
EELS
In-situ holder (heating , biasing, gas, liquid….)
….

## Slide 87: Transmission Electron Microscopy
Cold FEG: small coherent and intense beam: high spatial resolution for imaging and chemical analyses
Operating voltages: 200, 120 et 60 kV: large range of observable samples
Probe-corrected for high resolution in STEM -mode: 200kV= 70pm; 120kV: 90pm; 60kV: 110pm
High sensitivity EDX (detection solid angle: 1,8sr): high resolution chemical analyses and chemical maps
4D STEM with NANOMEGAS ASTAR module
CRHEA’s ACT-M SPECTRA 200 technical characteristics
https://www.crhea.cnrs.fr/ACT-M/
Ileana.florea@crhea.cnrs.fr

## Slide 88: Transmission Electron Microscopy
Scanning TEM
GaN
AlGaN1
AlGaN2
HAADF
Z contrast
ADF
Diffraction contrast
L. Lacomblez, P.M. Coulon (CRHEA); M. Cabié (IM2NP Marseille): FIB sample preparation
GaN pillar
20µm
2µm

## Slide 89: Transmission Electron Microscopy
Scanning TEM: HR-HAADF
GaN (0001)
Direct projection of the structure
(S)TEM image: 2D image of a 3D object

## Slide 90: Transmission Electron Microscopy
Scanning TEM: EDX

## Slide 91: Transmission Electron Microscopy
Scanning TEM: EDX
Quantitative EDX:
  Cliff-Lorimer method: CA/CB = k・IA/IB
Ga: 85%; Al:15%
Ga: 24%; Al:76%
Ga: 18%; Al:82%
EDX: Ga: 18%; Al: 82%
Strain from XRD: Ga: 17%; Al: 83%

## Slide 92: Transmission Electron Microscopy
Electron energy loss spectroscopy
Inelastic interaction: energy loss of e- => spectrum
Low-loss (< ≈ 50eV):
  Dielectric response
  Free electron density
  Band gap
  Specimen thickness determination
See M. Amato presentation for an example

## Slide 93: Transmission Electron Microscopy
Electron energy loss spectroscopy
Inelastic interaction: energy loss of e- => spectrum
High loss (core loss): characteristic edges
    Composition determination (low Z elements)
    Bonding/valence state
    Nearest-neighbor atomic structure
https://eels.info/about/overview

## Slide 94: Transmission Electron Microscopy
Electron energy loss spectroscopy
TiL12
LaM45
SrL23
MnL23
https://www.gatan.com/atomic-level-eels-mapping-using-high-energy-edges-dualeels-mode

## Slide 95: Transmission Electron Microscopy
HR-(S)TEM
Interface analysis
AlN
Si(111)
FFT highlighting dislocations
Distance beetwen misfit dislocations for a complete prelaxation complète:
D = dcds/(ds-dc)
AlN/Si(111): 5 AlN planes/ 4 Si planes: complete relaxation
N. Mante et al., J. Appl. Phys, 123, 715201(2018)

## Slide 96: Transmission Electron Microscopy
HR-(S)TEM
A-planeZnO on R-planesaph
Interface analysis

## Slide 97: Transmission Electron Microscopy
HR-(S)TEM
Strain analysis
GaN
AlGaN1
AlGaN2
AlGaN/AlGaN multi quantum well
geometrical Phase analysis (GPA) strain map

## Slide 98: Transmission Electron Microscopy
F.A. Ponce Group III Nitride Semiconductor Compounds, Physics and Applications (Clarendon Press, Oxford 1998).
III-N layers microstructure
Columnar microstructures (h>>w) dominated by twist
h
w
Plan view observation
AlN/Si MBE
S. Renesson, F. Semond (CRHEA)
Rotation map from geometrical Phase analysis (GPA)
HR-(S)TEM
Defect analysis

## Slide 99: Transmission Electron Microscopy
HR-(S)TEM
Defect analysis
Rotation map
8
9
8
9
8
8
9
9
b
b
a2
a1
Misorientation between subgrains compensated by the introduction of a-type edge dislocations

## Slide 100: Transmission Electron Microscopy
Scanning TEM: EDX
Atomic scale EDX
O
Sr
Fe
La
Ti
SrTiO3
SrLaFeO3
A. Polfelski, Brookhaven National Laboratory (US), Mc Master University(Canada)

## Slide 101: Transmission Electron Microscopy
D
C
A
B
Shifted
diffraction
pattern
1. I. Lazić, E.G.T. Bosch, Advances in Imaging and Electron Physics 199 (2017) 75-184.
2. E. Yücelen, I. Lazić, E.G.T. Bosch, Scientific Reports 8 (2018) 2676.
HAADF: I≈ Z2 => difficult to see both heavy and light atoms
Integrated differential phase imaging: iDPC
Imaging light elements

## Slide 102: Transmission Electron Microscopy
Imaging light elements
AlN
AlN
NbN
HAAADF image
AlN/NbN/AlN heterostructures
ZN=7; ZAl=13; ZNb=41
iDPC image
Al-polarity
N
Al
Nb
N-polarity
Al-polarity
N-polarity
A. Pédèches, H. Rotella, F. Semond (CRHEA)

## Slide 103: Transmission Electron Microscopy
4D STEM
Ophus C (2019) Microsc Microanal. doi:10.1017/S1431927619000497
Improvement of diffraction patterns with precession
Nanomegas solution

## Slide 104: Transmission Electron Microscopy
4D STEM

## Slide 105: Transmission Electron Microscopy
4D STEM
Applications:
    Comparison of experimental diffraction patterns with a library of calculated ones
      Phase mapping
      Orientation mapping
      Strain mapping at mesoscopic scale
εxx
N. Mante, PhD thesis Grenoble 2016

## Slide 106: Transmission Electron Microscopy
Sub Angstrom resolution => direct imaging of atomic columns: unsurpassed spatial resolution
Structural and chemical imaging
Localized information
Study in direct and reciprocal space
Requires time consuming sample preparation
Destructive

## Slide 107: Transmission Electron Microscopy
Basic
Advanced
Direct space | Reciprocal space
AFM; SEM SE imaging | XRD /2
SEM EDX/ECCI … | XRD RSM, φ scans ….
(S)TEM | XRD; (S)TEM
Don’t hesitate to contact specialists when you want to perform/understand advanced characterizations
Non destructive
destructive

## Slide 108
Merci pour votre attention!

## Slide 109
Atomic force microscopy
KPFM (Kelvin probe Force Microscopy)
Tapping mode
Simultaneous acquisition of a topographic image and an image of the work output
Information on local composition and surface electronic structure
By Sraisac - Own work, CC BY-SA 4.0, https://commons.wikimedia.org/w/index.php?curid=79006406
https://www.bruker.com/fr/products-and-solutions/microscopes/materials-afm/afm-modes/kpfm.html
Schematic diagram of the Kelvin method
KPFM: by applying a voltage between tip and sample, we can cancel out the tip/sample potential difference and determine the local difference in output work between tip and sample

## Slide 110
substrate
Scanning Electron Microscopy
substrate
Use a thin sample
(TEM or transmission in a SEM)
EDS in cross section

## Slide 111
Scanning Electron Microscopy
ECCI
A. Mandal et al., Microscopy and Microanalysis, 2024, 00, 1–14
MOVPE-grown GaN/Si

## Slide 112
Atomic force microscopy
Evolution of MoS2/GaN MBE growth as a function of time
rms: 0,32nm
rms: 0,13nm
rms: 0,26nm
rms: 0,2nm
30’
1h
2h
3h
Nucléation
Lateral triangular growth
coalescence
M. Al Khalfioui et al., J. of Cryst. Growth 652 (2025) 128047

## Slide 113
Scanning Electron Microscopy
µLED fabrication
EBSD

## Slide 114
X-Ray diffraction

2
Detector
Grazing incidence Xray diffraction; in plane diffraction: study of planes perpendicular to the surface